**Vetscan QR Testing**

**Table of Contents**

[1 Overview 2](#_Toc78370125)

[1.1 Scope of this Document 2](#_Toc78370126)

[1.2 Test Equipment 2](#_Toc78370127)

[1.3 Documentation 2](#_Toc78370128)

[1.4 QR Specification 3](#_Toc78370129)

[1.4.1 Content of QR code 3](#_Toc78370130)

[1.5 FUSE VETXML Test names 4](#_Toc78370131)

[1.6 Test Labels 5](#_Toc78370132)

[1.7 Test Label Generation 5](#_Toc78370133)

[1.8 Test Label Printing 6](#_Toc78370134)

[2 Test #1: 7](#_Toc78370135)

[2.1 Test Setup 7](#_Toc78370136)

[3 Test Instructions 8](#_Toc78370137)

[3.1 Web App 8](#_Toc78370138)

[3.2 Galaxy S10 8](#_Toc78370139)

[3.3 Python App 8](#_Toc78370140)

[4 Test Results 9](#_Toc78370141)

[4.1 Vetscan Hub Mockup Web App 9](#_Toc78370142)

[4.2 Galaxy S10 Cell Phone 10](#_Toc78370143)

[4.3 Laptop Web App 11](#_Toc78370144)

[4.4 Laptop Python App 12](#_Toc78370145)

[4.5 Vetscan Hub Mockup Python App 13](#_Toc78370146)

[5 Conclusions 14](#_Toc78370147)

[6 Revision History 15](#_Toc78370148)

# Overview

This document will contain a test plan and report involving the Vetscan’s camera. This will mainly focus on the Vetscan’s camera ability to read QR codes.

Each combination of label size, module size, and total characters will be tested. Each test will be recorded as pass or fail.

## Scope of this Document

This document will be used for engineering staff to plan and document testing.

## Test Equipment

* Vetscan Hub mockup unit
* Ruler used to measure distance from the camera to a QR code, and the size of the QR code label.
* Printed QR Labels. Encoded according to ISO 18004:2006, using Level M error correction, and the max amount of data that can be contained in a QR Label.
* One carboard box – 1.7 inches tall and wider than 50 mm.
* One carboard box – 3.5 inches tall and wider than 50 mm.
* qr\_code\_genertor.py – A Python application used to generate QR labels.

## Documentation

|  |  |
| --- | --- |
| ISO 18004:2006 | QR code specification |
| <https://www.the-qrcode-generator.com/>scan | Web site – create or read a QR label with given data |

## QR Specification

### Content of QR code

**Format:** the content of the QR code must be a valid json structure.

**Name:** Mandatory information. Valid names are current FUSE VETXML Test codes.

See Name table below.

Maximum length 8 characters.

**Lot:** Mandatory information. A string consistent with the lot id from the manufacturer.

**Exp:** Mandatory information. Expiration date expressed according to ISO 8601 as yyyy-mm-dd.

**Data:** Optional. String of data to be sent verbatim to the analyzer.

Valid QR code according to specification (ISO 18004:2006)

Version: Minimum 1 (21 × 21 modules)  
 Maximum 40 (177 × 177 modules)

Size: Minimum size 10 mm × 10 mm.  
 Maximum size 50 mm × 50 mm.

Error correction: Level M.

Printed Labels: printed on a high-contrast background

Example JSON string that would be encoded into a QR label.

{

"analyzer": "<analyzer id>",

"uuid": "<uuid>",

"lot": "<string>",

"exp": "<yyyy-mm-dd>",

"data": "<optional data>"

}

This is 53 characters long without any data or whitespace, just the name value pairs and syntax requirements. An example SPE QR would be as follows:

{

"analyzer":"vetscan-spe",

"exp":"2022-05-07",

"uuid":"0b7ec890-3960-11eb-a081-2790e47ff2f4",

"lot":"1234",

"data":""

}

This is 114 characters when whitespace is removed.

## FUSE VETXML Test names

|  |  |  |
| --- | --- | --- |
| **Analyzer** | **Name** | **Description** |
| Chemistry | AR | Avian/Reptilian Profile Plus |
| CDP | Comprehensive Diagnostic |
| CC | Critical Care Plus |
| EPP | Equine Profile Plus |
| KPP | Kidney Profile Plus |
| LA | Large Animal Profile |
| MLP | Mammalian Liver Profile |
| Prep | Prep Profile II |
| T4 | Thyroxine(T4)/Cholesterol Test |
| EP | Electrolyte Plus |
| PCP | Preventive Care Profile Plus |
| PP | Phenobarbital Profile |
| Hematology | HEM | Hematology |
| Chemistry2 | PT/aPTT | PT/aPTT Combination Test |
| Fib | Equine Fibrinogen Test |
| EFib | Equine Fibrinogen |
| CFib | Canine Fibrinogen |
| CBT | Canine Blood Typing |
| FBT | Feline Blood Typing |
| PTaPTT | PT/aPTT Combination |
| Phb | Phenobarbital (not released) |
| RapidTests | E | Ehrlichia Rapid Test |
| P | Parvo Rapid Test |
| A | Anaplasma Rapid Test |
| G | Giardia Rapid Test |
| L | Lyme Rapid Test |
| FF | FeLV\_Fiv Rapid Test |
| c | cPL Rapid Test |
| FLE | Flex4 Rapid Test |
| H | Heartworm Rapid Test |
| Fecalanalysis | FOVA | Fecal Ova/Oocysts |
| FGIA | Fecal Giardia |
| Urinalysis | SA | Urine Sediment |
| SA10 | Urine Sediment and UA10 |
| SA14 | Urine Sediment and UA14 |
| UA10 | UA10 |
| UA14 | UA14 |

## Test Labels

A QR code (abbreviated from Quick Response code) is a type of matrix barcode (or two-dimensional barcode[1]) invented in 1994 by the Japanese automotive company Denso Wave.[2] A barcode is a machine-readable optical label that contains information about the item to which it is attached. In practice, QR codes often contain data for a locator, identifier, or tracker that points to a website or application. A QR code uses four standardized encoding modes (numeric, alphanumeric, byte/binary, and kanji) to store data efficiently; extensions may also be used.

* The test labels will be printed on paper with matte and glossy finishes.
* The test labels will be scanned on boxes that are 1.7 inches and 3.5 inches tall.
* The test Label Sizes that will be used will be 10 mm x 10 mm and 50 x 50 mm.
* The QR codes will contain example JSON data which is a combination of Numeric, Alphanumeric and Binary characters. Therefor the encodeing of the JSON data will take more space then purly Numeric or Alphanumeric data.
  + **Numeric:**
    - **0 1 2 3 4 5 6 7 8 9**
  + **Alphanumeric:**
    - **0–9**
    - **A–Z**
    - **(upper-case only)**
    - **space**
    - **$ % \* + - . / :**
  + **Binary:**
    - **a-z (lower-case)**
    - **{ } \n “ ,**
    - **Every other character**
* For example, when using the Version 20 QR Code with correction level M, the maximum allowable characters are:
  + **Numeric: 1600**
  + **Alphanumeric: 970**
  + **Binary: 666**

## Test Label Generation

* The qr\_code\_generator.py Python application was used to generate the images for the different label data densities.
  + https://github.com/lincolnloop/python-qrcode
* The app created image files in the QR\_labels\_Alphanumeric\_JSON directory.

## Test Label Printing

Using the program ifranview (<https://www.irfanview.com/>), it is possible to scale and print all the QR codes at once.

1. Install irfanview
2. Open irfanview
3. Select File->Thumbnails
4. Using file view pane on left hand side of screen, navigate to the directory where the QR code images are located
5. Now click on the files to select the ones you wish to print
6. Select File->Print selected files as single images (batch print)
7. Select “Print Size:” “Custom”. Set the Width and Height.
   1. For the 10mm x 10mm QR codes: 1.1 cm
   2. For the 20mm x 20mm QR codes: 2.1 cm
   3. For the 50mm x 50mm QR codes: 5.3 cm
8. Now check the box for Headnote and put $D $F in the input box so that the file name is printed with the image
9. In the “Position on paper” section, select “Center horz”
10. In the “Position on paper” section, enter a “Top margin” of 5.00
11. You can go into the Printer setup and set the printer options on the printer driver (2 sided, flip, etc)
12. See below.



# Test #1:

The purpose of this test is to determine what is the maximum amount of data that can be read with the camera at given distances, label sizes, and label densities.

## Test Setup

* The web camera and QR labels will be tested using a web site that can activate the camera and record the QR label’s code. The web page will display the camera image and display the QR code.
* The QR labels will be placed on the top of ether the 1.7-inch or 3.5-inch-tall box.
* The box will be placed on the base of the Vetscan Hub Mockup.
* The display of the Vetscan Hub Mockup will need to be tilted until it aligns to the label.
* The web site will use the Vetscan Hub Mockup’s camera to capture an image of the QR label and attempt to decode the embedded data in the QR label.
* If the QR label is decoded, the data will be displayed on the web page. Each test case shall be recorded in the tables below. If the QR label is decoded, then the test case shall be marked as “Pass”, else it shall be marked as “Fail”.
* See screen capture of web page below.



# Test Instructions

* For 10 x 10 mm and 20 x 20 mm, use a 3.5” box to elevate the QR code off the base of the Vetscan Hub Mockup
* For 50 x 50 mm, use a 1.5” box to elevate the QR code off the base of the Vetscan Hub Mockup

## Web App

Perform the following test setup once at the beginning of testing:

1. Turn on the Vetscan Hub Mockup or Laptop that is under test.
2. Open the web site: <https://4qrcode.com/scan-qr-code.php>.
3. On the web page, click on the button labelled “Open camera”.
4. You should see an image from the unit’s camera displayed on the web page.

For each test case, select the proper box and QR label.

1. Place the test box under the Vetscan Hub Mockup’s or Laptop’s camera.
2. Place the QR label on the top of the box if performing test on Vetscan Hub Mockup.
3. Using the web page’s camera image, center the QR tag’s image in the center of the image.
4. If the QR label can be decoded, the web page will display the QR label’s data on web page.
5. Once the image is centered, wait at most 30 seconds for the web site to decode the QR label.
6. If the QR label was decoded within 30 seconds, then the test case passed, else the test case failed.
7. Record the pass/fail status of each test case.

## Galaxy S10

Perform the following test setup once at the beginning of testing:

1. Open the QR & Barcode Reader – TeaCapps on the Galaxy S10.

For each test case, select the proper box and QR label:

1. Place the test box under the Galaxy S10’s camera.
2. Using the camera view on the phone, center the QR tag’s image in the center of the image.
3. If the QR label can be decoded, the value will be displayed on the phone screen.
4. If the QR was decoded within 30 seconds, then the testcase passes, else the test case failed.
5. Record the pass/status of each test case.

## Python App

Perform the following test setup once at the beginning of testing:

1. Turn on the Vetscan Hub Mockup or the Laptop being tested.

For each test case, select the proper box and QR label.

1. Place the test box under the Vetscan Hub Mockup’s or Laptop’s camera.
2. Place the QR label on the top of the box if performing test on Vetscan Hub Mockup.
3. Initiate the qr scanner python app from the console.
4. Using the displayed camera image, center the QR tag’s image in the center of the image.
5. If the QR label can be decoded, the output will be displayed on the terminal.
6. Once the image is centered, wait at most 30 seconds for the web site to decode the QR label.
7. If the QR label was decoded within 30 seconds, then the test case passed, else the test case failed.
8. Record the pass/fail status of each test case.

## 

# Test Results

## Vetscan Hub Mockup Web App

* These test results were recorded on the Vetscan.
* Web page: <https://www.the-qrcode-generator.com/scan>
* Vetscan monitor titled to max angle.
* Box hieght: 10mm x 10mm: 9.5cm, 20mm x 20mm: 9cm, 50mm x 50mm: 9cm

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Version** | **Modules** | **Max**  **AlphaNum** | **Test**  **JSON** | **Pass/Fail**  **10 x 10 mm** | **Pass/Fail**  **20 x 20 mm** | **Pass/Fail**  **50 x 50 mm** |
| 1 | 21 x 21 | 20 | Too small | n/a | n/a | n/a |
| 2 | 25 x 25 | 38 | Too small | n/a | n/a | n/a |
| 3 | 29 x 29 | 61 | Too small | n/a | n/a | n/a |
| 4 | 33 x 33 | 90 | Too small | n/a | n/a | n/a |
| 5 | 37 x 37 | 122 | 88 | Fail | Pass | Pass |
| 6 | 41 x 41 | 154 | 120 |  | Pass | Pass |
| 7 | 45 x 45 | 178 | 144 |  | Pass | Pass |
| 8 | 49 x 49 | 221 | 187 |  | Pass | Pass |
| 9 | 53 x 53 | 262 | 228 |  | Pass | Pass |
| 10 | 57 x 57 | 311 | 274 |  | Pass | Pass |
| 11 | 61 x 61 | 366 | 329 |  | Pass | Pass |
| 12 | 65 x 65 | 419 | 382 |  | Pass | Pass |
| 13 | 69 x 69 | 483 | 446 |  | Fail | Pass |
| 14 | 73 x 73 | 528 | 491 |  | Fail | Pass |
| 15 | 77 x 77 | 600 | 564 |  | Fail | Fail |
| 16 | 81 x 81 | 656 | 619 |  |  | Fail |
| 17 | 85 x 85 | 734 | 697 |  |  |  |
| 18 | 89 x 89 | 816 | 779 |  |  |  |
| 19 | 93 x 93 | 909 | 872 |  |  |  |
| 20 | 97 x 97 | 970 | 933 |  |  |  |
| 21 | 101 x 101 | 1035 | 999 |  |  |  |
| 22 | 105 x 105 | 1134 | 1097 |  |  |  |
| 23 | 109 x 109 | 1248 | 1211 |  |  |  |
| 24 | 113 x 113 | 1326 | 1289 |  |  |  |
| 25 | 117 x 117 | 1451 | 1415 |  |  |  |
| 26 | 121 x 121 | 1542 | 1505 |  |  |  |
| 27 | 125 X 125 | 1637 | 1600 |  |  |  |
| 28 | 129 X 129 | 1732 | 1695 |  |  |  |
| 29 | 133 X 133 | 1839 | 1803 |  |  |  |
| 30 | 137 x 137 | 1994 | 1957 |  |  |  |
| 31 | 141 x 141 | 2113 | 2076 |  |  |  |
| 32 | 145 x 145 | 2238 | 2201 |  |  |  |
| 33 | 149 x 149 | 2369 | 2332 |  |  |  |
| 34 | 153 x 153 | 2506 | 2469 |  |  |  |
| 35 | 157 x 157 | 2632 | 2595 |  |  |  |
| 36 | 161 x 161 | 2780 | 2744 |  |  |  |

## Galaxy S10 Cell Phone

* These test results were recorded on a Galaxy S10 cell phone.
* Andriod app: QR & Barcode Reader - TeaCapps
* Distance from camera: ~2-inches

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Version** | **Modules** | **Max**  **AlphaNum** | **Test**  **JSON** | **Pass/Fail**  **10 x 10 mm** | **Pass/Fail**  **20 x 20 mm** | **Pass/Fail**  **50 x 50 mm** |
| 1 | 21 x 21 | 20 | Too small | n/a | n/a | n/a |
| 2 | 25 x 25 | 38 | Too small | n/a | n/a | n/a |
| 3 | 29 x 29 | 61 | Too small | n/a | n/a | n/a |
| 4 | 33 x 33 | 90 | Too small | n/a | n/a | n/a |
| 5 | 37 x 37 | 122 | 88 | Pass | Pass | Pass |
| 6 | 41 x 41 | 154 | 120 | Pass | Pass | Pass |
| 7 | 45 x 45 | 178 | 144 | Pass | Pass | Pass |
| 8 | 49 x 49 | 221 | 187 | Pass | Pass | Pass |
| 9 | 53 x 53 | 262 | 228 | Pass | Pass | Pass |
| 10 | 57 x 57 | 311 | 274 | Pass | Pass | Pass |
| 11 | 61 x 61 | 366 | 329 | Pass | Pass | Pass |
| 12 | 65 x 65 | 419 | 382 | Pass | Pass | Pass |
| 13 | 69 x 69 | 483 | 446 | Pass | Pass | Pass |
| 14 | 73 x 73 | 528 | 491 | Pass | Pass | Pass |
| 15 | 77 x 77 | 600 | 564 | Fail | Pass | Pass |
| 16 | 81 x 81 | 656 | 619 |  | Pass | Pass |
| 17 | 85 x 85 | 734 | 697 |  | Pass | Pass |
| 18 | 89 x 89 | 816 | 779 |  | Pass | Pass |
| 19 | 93 x 93 | 909 | 872 |  | Pass | Pass |
| 20 | 97 x 97 | 970 | 933 |  | Pass | Pass |
| 21 | 101 x 101 | 1035 | 999 |  | Pass | Pass |
| 22 | 105 x 105 | 1134 | 1097 |  | Pass | Pass |
| 23 | 109 x 109 | 1248 | 1211 |  | Pass | Pass |
| 24 | 113 x 113 | 1326 | 1289 |  | Pass | Pass |
| 25 | 117 x 117 | 1451 | 1415 |  | Pass | Fail |
| 26 | 121 x 121 | 1542 | 1505 |  | Pass |  |
| 27 | 125 X 125 | 1637 | 1600 |  | Pass |  |
| 28 | 129 X 129 | 1732 | 1695 |  | Pass |  |
| 29 | 133 X 133 | 1839 | 1803 |  | Pass |  |
| 30 | 137 x 137 | 1994 | 1957 |  | Fail |  |
| 31 | 141 x 141 | 2113 | 2076 |  | Fail |  |
| 32 | 145 x 145 | 2238 | 2201 |  | Fail |  |
| 33 | 149 x 149 | 2369 | 2332 |  |  |  |
| 34 | 153 x 153 | 2506 | 2469 |  |  |  |
| 35 | 157 x 157 | 2632 | 2595 |  |  |  |
| 36 | 161 x 161 | 2780 | 2744 |  |  |  |
| 37 | 165 x 165 | 2894 | 2857 |  |  |  |
| 38 | 169 x 169 | 3054 | 3017 |  |  |  |

## Laptop Web App

* These test results were recorded on a ThinkPad P50.
* Web page: https://www.the-qrcode-generator.com
* Distance from Web camera: ~2.25-inches

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Version** | **Modules** | **Max**  **AlphaNum** | **Test**  **JSON** | **Pass/Fail**  **10 x 10 mm** | **Pass/Fail**  **20 x 20 mm** | **Pass/Fail**  **50 x 50 mm** |
| 1 | 21 x 21 | 20 | Too small | n/a | n/a | n/a |
| 2 | 25 x 25 | 38 | Too small | n/a | n/a | n/a |
| 3 | 29 x 29 | 61 | Too small | n/a | n/a | n/a |
| 4 | 33 x 33 | 90 | Too small | n/a | n/a | n/a |
| 5 | 37 x 37 | 122 | 88 | Fail | Pass | Pass |
| 6 | 41 x 41 | 154 | 120 |  | Pass | Pass |
| 7 | 45 x 45 | 178 | 144 |  | Pass | Pass |
| 8 | 49 x 49 | 221 | 187 |  | Pass | Pass |
| 9 | 53 x 53 | 262 | 228 |  | Pass | Pass |
| 10 | 57 x 57 | 311 | 274 |  | Fail | Pass |
| 11 | 61 x 61 | 366 | 329 |  |  | Pass |
| 12 | 65 x 65 | 419 | 382 |  |  | Pass |
| 13 | 69 x 69 | 483 | 446 |  |  | Pass |
| 14 | 73 x 73 | 528 | 491 |  |  | Pass |
| 15 | 77 x 77 | 600 | 564 |  |  | Pass |
| 16 | 81 x 81 | 656 | 619 |  |  | Pass |
| 17 | 85 x 85 | 734 | 697 |  |  | Fail |
| 18 | 89 x 89 | 816 | 779 |  |  |  |
| 19 | 93 x 93 | 909 | 872 |  |  |  |
| 20 | 97 x 97 | 970 | 933 |  |  |  |
| 21 | 101 x 101 | 1035 | 999 |  |  |  |
| 22 | 105 x 105 | 1134 | 1097 |  |  |  |
| 23 | 109 x 109 | 1248 | 1211 |  |  |  |
| 24 | 113 x 113 | 1326 | 1289 |  |  |  |
| 25 | 117 x 117 | 1451 | 1415 |  |  |  |
| 26 | 121 x 121 | 1542 | 1505 |  |  |  |
| 27 | 125 X 125 | 1637 | 1600 |  |  |  |
| 28 | 129 X 129 | 1732 | 1695 |  |  |  |
| 29 | 133 X 133 | 1839 | 1803 |  |  |  |
| 30 | 137 x 137 | 1994 | 1957 |  |  |  |
| 31 | 141 x 141 | 2113 | 2076 |  |  |  |
| 32 | 145 x 145 | 2238 | 2201 |  |  |  |
| 33 | 149 x 149 | 2369 | 2332 |  |  |  |
| 34 | 153 x 153 | 2506 | 2469 |  |  |  |
| 35 | 157 x 157 | 2632 | 2595 |  |  |  |
| 36 | 161 x 161 | 2780 | 2744 |  |  |  |

## Laptop Python App

* These test results were recorded on a ThinkPad P50.
* Python app: qr\_scanner.py
* Distance from Web camera: ~2.25-inches

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Version** | **Modules** | **Max**  **AlphaNum** | **Test**  **JSON** | **Pass/Fail**  **10 x 10 mm** | **Pass/Fail**  **20 x 20 mm** | **Pass/Fail**  **50 x 50 mm** |
| 1 | 21 x 21 | 20 | Too small | n/a | n/a | n/a |
| 2 | 25 x 25 | 38 | Too small | n/a | n/a | n/a |
| 3 | 29 x 29 | 61 | Too small | n/a | n/a | n/a |
| 4 | 33 x 33 | 90 | Too small | n/a | n/a | n/a |
| 5 | 37 x 37 | 122 | 88 | Fail | Pass | Pass |
| 6 | 41 x 41 | 154 | 120 |  | Pass | Pass |
| 7 | 45 x 45 | 178 | 144 |  | Pass | Pass |
| 8 | 49 x 49 | 221 | 187 |  | Fail | Pass |
| 9 | 53 x 53 | 262 | 228 |  |  | Pass |
| 10 | 57 x 57 | 311 | 274 |  |  | Pass |
| 11 | 61 x 61 | 366 | 329 |  |  | Pass |
| 12 | 65 x 65 | 419 | 382 |  |  | Pass |
| 13 | 69 x 69 | 483 | 446 |  |  | Pass |
| 14 | 73 x 73 | 528 | 491 |  |  | Pass |
| 15 | 77 x 77 | 600 | 564 |  |  | Pass |
| 16 | 81 x 81 | 656 | 619 |  |  | Pass |
| 17 | 85 x 85 | 734 | 697 |  |  | Pass |
| 18 | 89 x 89 | 816 | 779 |  |  | Pass |
| 19 | 93 x 93 | 909 | 872 |  |  | Pass |
| 20 | 97 x 97 | 970 | 933 |  |  | Pass |
| 21 | 101 x 101 | 1035 | 999 |  |  | Pass |
| 22 | 105 x 105 | 1134 | 1097 |  |  | Pass |
| 23 | 109 x 109 | 1248 | 1211 |  |  | Pass |
| 24 | 113 x 113 | 1326 | 1289 |  |  | Pass |
| 25 | 117 x 117 | 1451 | 1415 |  |  | Pass |
| 26 | 121 x 121 | 1542 | 1505 |  |  | Fail |
| 27 | 125 X 125 | 1637 | 1600 |  |  |  |
| 28 | 129 X 129 | 1732 | 1695 |  |  |  |
| 29 | 133 X 133 | 1839 | 1803 |  |  |  |
| 30 | 137 x 137 | 1994 | 1957 |  |  |  |
| 31 | 141 x 141 | 2113 | 2076 |  |  |  |
| 32 | 145 x 145 | 2238 | 2201 |  |  |  |
| 33 | 149 x 149 | 2369 | 2332 |  |  |  |
| 34 | 153 x 153 | 2506 | 2469 |  |  |  |
| 35 | 157 x 157 | 2632 | 2595 |  |  |  |
| 36 | 161 x 161 | 2780 | 2744 |  |  |  |

## Vetscan Hub Mockup Python App

* These test results were recorded on the Vetscan.
* Python app: qr\_scanner.py
* Box hieght: 10mm x 10mm: 9.5cm, 20mm x 20mm: 9cm, 50mm x 50mm: 9cm

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| **Version** | **Modules** | **Max**  **AlphaNum** | **Test**  **JSON** | **Pass/Fail**  **10 x 10 mm** | **Pass/Fail**  **20 x 20 mm** | **Pass/Fail**  **50 x 50 mm** |
| 1 | 21 x 21 | 20 | Too small | n/a | n/a | n/a |
| 2 | 25 x 25 | 38 | Too small | n/a | n/a | n/a |
| 3 | 29 x 29 | 61 | Too small | n/a | n/a | n/a |
| 4 | 33 x 33 | 90 | Too small | n/a | n/a | n/a |
| 5 | 37 x 37 | 122 | 88 | Fail | Pass | Pass |
| 6 | 41 x 41 | 154 | 120 |  | Pass | Pass |
| 7 | 45 x 45 | 178 | 144 |  | Pass | Pass |
| 8 | 49 x 49 | 221 | 187 |  | Pass | Pass |
| 9 | 53 x 53 | 262 | 228 |  | Pass | Pass |
| 10 | 57 x 57 | 311 | 274 |  | Pass | Pass |
| 11 | 61 x 61 | 366 | 329 |  | Pass | Pass |
| 12 | 65 x 65 | 419 | 382 |  | Pass | Pass |
| 13 | 69 x 69 | 483 | 446 |  | Fail | Pass |
| 14 | 73 x 73 | 528 | 491 |  |  | Pass |
| 15 | 77 x 77 | 600 | 564 |  |  | Pass |
| 16 | 81 x 81 | 656 | 619 |  |  | Pass |
| 17 | 85 x 85 | 734 | 697 |  |  | Pass |
| 18 | 89 x 89 | 816 | 779 |  |  | Pass |
| 19 | 93 x 93 | 909 | 872 |  |  | Pass |
| 20 | 97 x 97 | 970 | 933 |  |  | Pass |
| 21 | 101 x 101 | 1035 | 999 |  |  | Pass |
| 22 | 105 x 105 | 1134 | 1097 |  |  | Pass |
| 23 | 109 x 109 | 1248 | 1211 |  |  | Pass |
| 24 | 113 x 113 | 1326 | 1289 |  |  | Pass |
| 25 | 117 x 117 | 1451 | 1415 |  |  | Pass |
| 26 | 121 x 121 | 1542 | 1505 |  |  | Pass |
| 27 | 125 X 125 | 1637 | 1600 |  |  | Pass |
| 28 | 129 X 129 | 1732 | 1695 |  |  | Pass |
| 29 | 133 X 133 | 1839 | 1803 |  |  | Pass |
| 30 | 137 x 137 | 1994 | 1957 |  |  | Pass |
| 31 | 141 x 141 | 2113 | 2076 |  |  | Pass |
| 32 | 145 x 145 | 2238 | 2201 |  |  | Pass |
| 33 | 149 x 149 | 2369 | 2332 |  |  | Pass |
| 34 | 153 x 153 | 2506 | 2469 |  |  | Pass |
| 35 | 157 x 157 | 2632 | 2595 |  |  | Pass |
| 36 | 161 x 161 | 2780 | 2744 |  |  | Pass |
| 37 | 165 x 165 | 2894 | 2857 |  |  | Pass |
| 38 | 169 x 169 | 3054 | 3017 |  |  | Pass |
| 39 | 173 x 173 | 3220 | 3183 |  |  | Pass |
| 40 | 177 x 177 | 3391 | 3355 |  |  | Fail |

# Conclusions

The glare from the Vetscan’s camera light is causing glare in the captured image. The QR reader is not able to read the label when the glare appears within the image of the QR label.

The current version of the Vetscan’s camera does not have the high resolution needed to read 10mm x 10mm QR labels.

The current version of the Vetscan’s camera limits it to reading 20mm x 20mm QR labels that have about 382 characters using Version 12 encoding.

The current version of the Vetscan’s camera limits it to reading 50mm x 50mm QR labels that have about 3183 characters using Version 39 encoding.

The Vetscan’s camera focal length of the camera is so far from the lens that the high density, 10mm x 10mm QR labels can’t be read. At about 1 inch from the camera, the image is very small.

On the Vetscan, the Python app was much better at decoding the lables then the web app was.

# Revision History

Record the results of testing labels at 3 inches (76.2 mm).

|  |  |  |  |
| --- | --- | --- | --- |
| **Version** | **Date** | **Author** | **Comments** |
| A beta | 15 JUL 2021 | Bruce Graham | Initial work. The use of the 10 mm x 10 mm QR labels is to be resolved with initial testing. |
| 1 | 22 JUL 2021 | Brian Newberry | 1.4.1 Updated Content of label code. |
| 2 | 26 JUL 2021 | Bruce Graham | 4.5 Added test results Vetscan: 20mm  4.6 Added test results Laptop: 20mm -web app  4.7 Added test results Laptop: 20mm – Python app  4.8 Added test results Galaxy S10 cell phone: 20mm  4.9 Added test results – Vetscan: 20mm with Python app |
| 3 | 28 JUL 2021 | Bruce Graham | Condensed tables for each device. |